SOP - 008

PROCEDURE FOR

CHIKUNGUNYA IGG/IGM  RAPID TEST

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	9

15.	Interfering Substances	9

16.	Result Calculation	10

17.	Biological Reference Interval	10

18.	Reportable Interval	10

19.	Procedure when Results are not in Range	10

20.	Critical Alert Values	10

21.	Laboratory Clinical Interpretation	10

22.	Potential Source of Variability	11

23.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | CDC | Centre for Disease Control |
| 2 | EDTA | Ethylenediaminetetraacetic Acid |
| 3 | ELISA | Enzyme-Linked Immunosorbent Assay |
| 4 | EQAS | External Quality Assurance Scheme |
| 5 | IgG | Immunoglobulin G |
| 6 | IgM | Immunoglobulin M |
| 7 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 8 | SE | Serology |
| 9 | SOP | Standard Operating Procedure |
| 10 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - This SOP describes the procedure to test Chikungunya IgG/IgM antibodies by the rapid immunochromatography method.

- Scope

This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Chikungunya IgG/IgM rapid tests are quick diagnostic tools that detect antibodies (IgG and IgM) against the Chikungunya virus in a patient's blood. These tests can differentiate between a current (IgM positive) and past (IgG positive) infection, aiding in early diagnosis and management.

- Test Principle

Monoclonal antihuman IgG and IgM are immobilized on a porous nitrocellulose membrane at individual test lines – G and M respectively. Inactivated Chikungunya virus in the antigen pad and monoclonal anti Chikungunya E1 gold in the conjugate pad release by adding assay diluent and react with Chikungunya IgG/IgM antibodies if present in the Sample which further bind to immobilised antihuman IgG/IgM antibodies giving a distinct pinkish purple band against a white background at the test region – IgG/IgM respectively. At the control region (‘C’) a “Built-in Quality Control band” has been devised to confirm the proper functioning of the device, reagent and correct procedural application.

- Specific Performance Characteristics

| Sample panel | Chikungunya IgG | Chikungunya IgM |
| --- | --- | --- |
| Sensitivity | 100% | 100% |
| Specificity | 97.68% | 99.61% |

- Type of Sample

- Serum/Plasma or Whole Blood must be used for testing. Should a delay in testing occur, store the sample at 2-8°C. Samples can be stored for upto a week.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel tube / EDTA vacutainer.

- Equipment & Reagents

Test Device

Timer

Buffer

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Centrifuge calibration done on an annual basis by a NABL accredited calibration agency.

- Test Procedure

- Allow the test device and specimen to equilibrate to room temperature (15 - 300c) prior to testing.

Add 10 µl serum or plasma or whole blood to the sample region.

Add 3 drops of Buffer Solution near the buffer region.

READ RESULT after 15-30 minutes.

- IMPORTANT: It is important to allow each solution to soak in the test device before adding the next solution.

- Quality Control Procedure

Each testing device has an inbuilt Control Band to confirm the validity of the device. The appearance of the control dot after completing the procedural test indicates a valid test.

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – Interlaboratory comparison.

- Interfering Substances

- Kit works best when used with fresh samples, however the frozen or viscous samples can also perform well if the following instructions are strictly adhered to

Frozen Samples

      - Allow the sample to thaw in a vertical position in the rack. Mix the sample thoroughly. If particles are seen, allow them to settle at the bottom or if a centrifuge is available, the sample can be centrifuged at 3500 rpm. for 15 minutes.

      - Insert the dropper just below the top surface of the sample and withdraw one drop of the sample.

Thick or Viscous Samples

      - Whenever possible, a clear specimen should be used. However, viscous, thick or turbid samples which may sometimes take more than 40-60 seconds to flow through the membrane should be centrifuged at 3500 rpm. for 15 minutes and retested on a fresh device to avoid inconsistent results.

- Result Calculation

NEGATIVE

- The control line (C) is only visible on the test device. No IgG and IgM antibodies were detected.

IgM POSITIVE

- The control line (C) and IgM line (M) are visible on the test device.

IgM and IgG POSITIVE

- The control line (C), IgM line (M) and IgG are visible on the test device.

- IgG Positive Result: The control line (C) and IgG are visible on the test device.

INVALID

- The test is considered invalid if the control dot is not formed. Test should be repeated with new device.

- Biological Reference Interval

- Non-Reactive

- Reportable Interval

- 4 Hours

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- Chikungunya disease is an arboviral disease caused by Chikungunya Virus, transmitted by bite of Aedes aeqypti. The presentation of the disease may be asymptomatic with subclinical infection or symptomatic with fever, body pains, joint pains. Results with immunochromatography have to be confirmed by ELISA method or by Chikungunya Virus detection by molecular methods.

- Potential Source of Variability

A negative result cannot exclude a recent Chikungunya Virus infection.

A negative result can occur if the quantity of antibody present is below the detection limits. Results should be correlated clinically and with other laboratory investigations.

Optimum test performance depends on strict adherence to the test procedure as described in this manual. Any deviation from test procedure may lead to erratic results.

- References

- Cruikshank R, (1982), Medical Microbiology, 12th Edition, 403

- Chikungunya Test – SD BIOSENSOR Kit Insert