SOP - 010

PROCEDURE FOR

ANTI NUCLEAR ANTIBODY DETECTION BY LINE PROBE IMMUNOASSAY

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Patient Preparation	8

9.	Type of Container & Additives	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Test Procedure	9

14.	Quality Control Procedure	10

15.	Interfering Substances	10

16.	Result Calculation	11

17.	Biological Reference Interval	11

18.	Reportable Interval	11

19.	Procedure when Results are not in Range	11

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	12

22.	Potential Source of Variability	13

23.	References	13

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | Quality Manager | Dr. Sai Varun Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | ANA | Anti Nuclear Antibody |
| 2 | EDTA | Ethylenediaminetetraacetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | IVD | In Vitro Diagnostics |
| 5 | IgG | Immunoglobulin G |
| 6 | LIA | Line Probe Immunoassay |
| 7 | NABL | National Accreditation Board for Testing and Calibration Laboratories |
| 8 | SARD | Systemic Auto Immune Rheumatic Diseases |
| 9 | SE | Serology |
| 10 | SLE | Systemic Lupus Erythematosus |
| 11 | SOP | Standard Operating Procedure |
| 12 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

  - This SOP describes the procedure to detect Anti Nuclear Antibodies by the Line Probe Immunoassay method.

- Scope

This procedure applies to all staff of the Serology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Serology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

Anti Nuclear Antibody Line Probe Immunoassay is an IVD assay for the qualitative determination of 18 IgG class antibodies against dsDNA, Nucleosome, Histones, SmD1, ribosomal PO (RPP). SS-A/Ro 60, SS-A/Ro 52, SS-B/La, CENP-B, Sci70, U1-snRNP, AMA M2, Jo-1, PM-Scl, Mi-2, Ku and DFS70 in human serum or plasma. The assay is intended for professional use as aid to diagnosis and classification of systemic auto immune rheumatic diseases (SARD).

- Test Principle

The test is based on the principle of the line immunoassay (LIA). Nuclear and associated cytosolic antigens are applied as lines on a nitrocellulose membrane. The nitrocellulose membrane is blocked to prevent unspecific reactions. Autoantibodies bind to antigens on the strip and a secondary horse radish peroxidase labelled Anti human IgG antibody is used which appears as brown line if antibodies against respective antigen are present.

- Specific Performance Characteristics

- Sera from antibody positive patients in IIF were tested to study the diagnostic sensitivity of the IMTEC-ANA-LIA XL. The diagnostic specificity was determined using serum samples from apparently healthy blood donors.

| Sensitivity | 99.1% |
| --- | --- |
| Specificity | 98% |

- Type of Sample

- Serum or Plasma must be used for testing. Should a delay in testing occur, store the sample at 2-8°C for upto a week or for one year at -20°C. Do not use hemolysed, icteric, lipemic serum.

- Patient Preparation

- No special preparation of the patient is required prior to specimen collection by approved techniques.

- Type of Container & Additives

- Collect 2 ml of venous blood in a plain red topped/ yellow gel tube / EDTA vacutainer.

- Equipment & Reagents

Wash buffer Solution (20Xx – Preparation - 1 part in 19 parts of distilled water)

Diluent buffer – ready to use

Conjugate solution

Substrate solution – ready to use

Test strips coated with antigen

Centrifuge to separate serum or plasma

HUMASCAN FA

Rotary shaker

Pipette – 1ml, 10µL

- Safety

The use of disposable gloves and proper biohazardous clothing is STRONGLY RECOMMENDED while running the test.

In case there is a cut or wound on the hand, DO NOT PERFORM THE TEST.

Do not smoke, drink or eat in areas where specimens or kit reagents are being handled.

Tests are for in vitro diagnostic use only and should be run by competent personnel only.

Do not pipette by mouth.

All materials used in the assay and samples should be decontaminated in 1% sodium hypochlorite solution for 30-60 min. They should be disposed of in accordance with established safety procedures.

Wash hands thoroughly with soap or any suitable detergent, after the use of the kit. Consult a physician immediately in case of accident or contact with eyes, in the event that contaminated material is ingested or comes in contact with skin punctures or wounds.

Spills should be decontaminated promptly with Sodium Hypochlorite or any other suitable disinfectant.

- Calibration Procedure

- Centrifuge calibration is done on an annual basis by a NABL accredited calibration agency.

- Test Procedure

Precautions

      - Do not interchange the reagents with other test kits and lots.

      - Any crystallised salt inside the wash buffer should be dissolved before use.

      - Do not dry the strip during incubation steps.

      - Do not touch with fingers, use tweezers.

      - Reagents to be stored at 2-8°C when not in use. Reagents stability is maintained up to 6 weeks as per kit insert after opening.

      - Do not use kit components beyond the expiration date.

      - Bring all the reagents and specimens to room temperature (25°C-30°C) before beginning the test. The immunological sequence of reactions which take place during different procedural steps shows best performance at room temperature.

Procedure

      - Place the required number of strips in the incubation tray at the working area.

      - Write the sample number to be tested on the incubation tray.

      - Sample preparation: Dilute 10µL of sample should be diluted with 1 ml of dilution buffer.

      - Place the strip into the incubation tray. Wet the membrane with diluted wash buffer and incubate for 1 minute at room temperature.

      - Remove the wash buffer and add diluted samples. Incubate for 30 minutes at room temperature.

      - Wash with wash buffer, incubate for 5 minutes with gentle agitation. Remove the wash and repeat it 3 times.

      - Add Conjugate and incubate for 30 minutes at room temperature. Repeat the wash step f.

      - Add substrate and incubate for 10 minutes at room temperature. Remove substrate.

      - Add distilled water and incubate for 1 minute at room temperature. Remove distilled water.

      - Add stop solution and incubate for 5 minutes at room temperature. Remove stop solution.

      - Dry the strip thoroughly and READ THE RESULTS IMMEDIATELY through HUMASCAN Reader.

- Quality Control Procedure

Each testing device has an inbuilt Control Band to confirm the validity of the device as well as a Cut-off Control Band. Intensity of functional control should be more than the intensity of the cut-off control.

All kits and lots will be tested as per reagent acceptability criteria.

EQAS participation – Interlaboratory comparison done annually with two samples.

- Interfering Substances

Frozen Samples

      - The ANA-LIA XL Test is best when used with fresh samples that have not been frozen and thawed. However, most frozen samples will perform well if the following suggested procedure is followed. Allow the sample to thaw in a vertical position in the rack. Do not shake the sample. This allows particles to settle to the bottom. If a centrifuge is available, the sample can be centrifuged at 3000 rpm for 15 minutes.

Thick or Viscous Samples

      - Whenever possible, a clear specimen should be used. However, viscous, thick or turbid samples which may sometimes take more than 40-60 seconds to flow through the membrane should be centrifuged at 3000 rpm for 15 minutes and retested on a fresh device to avoid inconsistent results.

- Result Calculation

Fix the strip onto the scoring sheet and align the reference line of the strip with the reference line on the scoring sheet.

Align the dotted reference line of the evaluation of the template with the reference line of the strip.

The interpretation of the test results takes place exclusively on the basis of the respective cutoff control regarded for each strip.

NEGATIVE

- The test result is negative, if no band seen or if the intensity is less in comparison to the cut off control.

EQUIVOCAL

- The test is equivocal, if the intensity of the test band and the cut off control band do not significantly differ.

POSITIVE

- The test result is positive, if the intensity of the test band is more in comparison to the cut-off control.

- Biological Reference Interval

- Non-Reactive

- Reportable Interval

- 24 Hours

- Procedure when Results are not in Range

- The test is repeated when the control band is not matching the acceptability criteria.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- ANA profile provides a comprehensive immunochromatographic assay for human autoantibodies of the IgG class to 14 different autoantigens diagnostically implicated in a varied range of auto immune diseases. The specificity for all specified autoantibodies has been found to be 100% in documented cohort studies, however the sensitivities are variable and may be correlated to the prevalence in particular diseases as mentioned below.

| Antibodies | Disease | Prevalence of auto antibodies |
| --- | --- | --- |
| U1-nRNP | Mixed Connective tissue disease | 95-100 % |
|  | Systemic Lupus Erythematosus | 15-40% |
|  | Systemic Sclerosis | 2-12% |
| Sm (Smith antigen) | Systemic Lupus Erythematosus | 5-40% |
|  | Systemic Sclerosis | 40-95% |
|  | Neonatal lupus | 95-100% |
| SS-A Ro 60kD | Sjogren syndrome | 40-80% |
|  | Disseminated Lupus Erythematosus | 30-40% |
|  | Primary biliary cirrhosis | 20% |
| SS-A Ro 52kD | Sjogren syndrome |  |
| SS-B(La) | Systemic Sclerosis | 40-95% |
| Scl-70 | Systemic Lupus Erythematosus | 20-60% |
| PM-Scl | Neonatal lupus | 95-100% |
|  | Systemic Sclerosis | 5-75% |
|  | Polymyositis-SS Overlap syndrome | 18% |
| Jo-1 | Polymyositis / Dermatomyositis | 25-35% |
| Centromeres | Systemic Sclerosis | 8-95% |
|  | PBS | 10-30% |
| PCNA (Proliferative cell nuclear antigen) | Systemic Lupus Erythematosus | 3% |
| ds DNA | Systemic Lupus Erythematosus | 40-90% |
| Nucleosomes | Systemic Lupus Erythematosus | 40-70% |
| Histones | Diffuse Lupus Erythematosis | 95-100% |
|  | Systemic Lupus Erythematosus | 50% |
|  | RA | 15-50% |
| Ribosomes P-protein | Systemic Lupus Erythematosus | 10% |
| Ku | Polymyositis | 5-25% |
| Mi-2 | Dermatomyositis | 95% |
| AMA-M2 | Primary Biliary Cirrhosis | >95% |
| DFS 70 | Not associated with systemic autoimmune rheumatic disease |  |
| CENP B (Centromere Protein B) | CREST syndrome | 70-90% |

| Reference | Cut-off Range | Result |
| --- | --- | --- |
| Negative | 0 – 0.79 | - |
| Borderline / Equivocal | 0.80 – 1.14 | 0 |
| Weak Positive | 1.15 – 2.49 | + |
| Mid-level Positive | 2.50 – 3.99 | ++ |
| Strong Positive | ≥ 4.00 | +++ |

- Potential Source of Variability

The intensity of the band does not correlate with the antibody titres.

Elevated levels of immunocomplexes may cause false positive results by non-specific binding.

Optimum test performance depends on strict adherence to the test procedure as described in this manual. Any deviation from test procedure may lead to erratic results.

This is only a screening test. Therefore, for a definitive diagnosis, the patient’s clinical history, symptomatology as well as serological data, should be considered.

- References

- ANA LIA – Kit Insert