SOP – 001

PROCEDURE FOR

BIOSAFETY CABINET

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Principle	7

5.	Required Equipment’s & Reagents	7

6.	Environment & Safety Control	8

7.	Calibration	9

8.	Equipment Record	9

9.	Procedure	10

10.	Maintenance	11

11.	Troubleshooting	11

12.	References	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | BSC | Biosafety Cabinet |
| 2 | °C | Degree Centigrade |
| 3 | HEPA | High Efficiency Particulate Air |
| 4 | HOD | Head of Department |
| 5 | MSP | Management System Procedure |
| 6 | N/A | Not Applicable/Available |
| 7 | PPE | Personal Protection Equipment |
| 8 | QC | Quality Control |
| 9 | QM | Quality Manager |
| 10 | SOP | Standard Operating Procedure |
| 11 | TDPL | TRUSTlab Diagnostics Private Limited |
| 12 | ULPA | Ultra-Low Particulate Air |
| 13 | UV | Ultraviolet |

- Objective

- The purpose of this SOP is to provide guidelines for operating and maintaining the Biosafety Cabinets at TRUSTlab’s NRL, to ensure safe and proper usage.

- Scope

- This procedure applies to all staff of the Molecular Biology Department.

- Responsibilities

It is the responsibility of the technicians of the respective Departments to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff that may be involved in this process (including contract personnel) will be responsible for complying with this procedure.

- Principle

- The primary purpose of a biosafety cabinet (BSC or Biological Safety Cabinet or Microbiological Safety Cabinet) is to protect the laboratory users and the environment from pathogens. BSC provides an enclosed and ventilated workspace. The enclosed space enhances safety when working with materials contaminated (or potentially contaminated) with pathogens. Biosafety cabinets find use where a defined biosafety level is required. Class II biosafety cabinets provide protection of the samples and of the environment, since makeup air is also HEPA-filtered. Class II biosafety cabinets are characterized by a limited fixed inward airflow access opening providing personnel protection, a vertical downward HEPA filtered work zone offering product protection and HEPA filtered exhaust providing environmental protection.

- Required Equipment’s & Reagents

      - 70% ethanol

      - Tissue

      - PPE

      - Disposable gloves

      - Eye Protection

      - Clean water & detergent

      - Sodium Hypochlorite

- Environment & Safety Control

| Safety Precautions | Avoid the use of flammable gases or solvents in the Biosafety Cabinet. An open flame should NOT be used in the Biosafety Cabinet. Open flames will disrupt airflow patterns, burn the HEPA filter and/or damage the filter’s adhesive. Gases under high pressure should not be used in the Biosafety Cabinet, as they may disrupt its airflow patterns. An open flame should NOT be used in the Biosafety Cabinet. An open flame should NOT be used in the Biosafety Cabinet. An open flame should NOT be used in the Biosafety Cabinet. An open flame should NOT be used in the Biosafety Cabinet. An open flame should NOT be used in the Biosafety Cabinet. Open flames will disrupt airflow patterns, burn the HEPA filter and/or damage the filter’s adhesive. Gases under high pressure should not be used in the Biosafety Cabinet, as they may disrupt its airflow patterns. An open flame should NOT be used in the Biosafety Cabinet. Open flames will disrupt airflow patterns, burn the HEPA filter and/or damage the filter’s adhesive. Gases under high pressure should not be used in the Biosafety Cabinet, as they may disrupt its airflow patterns. An open flame should NOT be used in the Biosafety Cabinet. Open flames will disrupt airflow patterns, burn the HEPA filter and/or damage the filter’s adhesive. Gases under high pressure should not be used in the Biosafety Cabinet, as they may disrupt its airflow patterns. Avoid direct exposure of plastic or coated materials to ultraviolet (UV) radiation. Avoid direct exposure of plastic or coated materials to ultraviolet (UV) radiation. Avoid splashing the disinfecting solution on skin or clothing. Ensure adequate ventilation. |
| --- | --- |

- Calibration

Frequency: Once a year

The complete range of documentation and testing services to ensure that your biological safety cabinetry provides the necessary clean air and containment for your work processes and your personnel.

      - HEPA/ULPA Filter Installation Leak Test

      - Airflow Velocity Profile Test

      - Airflow Smoke Pattern Test

      - Downflow Airflow Velocity and Volume Test

      - Inflow Airflow Velocity and Volume Test

      - Decontamination with Para formaldehyde

      - Particle Counting Test

      - Noise Level Test

      - Light Intensity Test

      - Filter Pressure Drop Test

      - UV Light Test

In addition, the following secondary maintenance services:

      - Supply and Installation of Pre-filters and HEPA Filters

      - Supply and Installation of Blower

      - Repair of Stainless Steel

      - Electrical/Mechanical Safety Checks

      - Repair of Front Hinged Window or Sash

- Equipment Record

      - Equipment Name: Biosafety Cabinet

      - Manufacturer Name: ICLEAN (Integrated Cleanroom Technologies Pvt. Ltd.)

      - Serial No & equipment ID:

        - 2021012 & TDPL/PR/CL-039

        - 2021013 & TDPL/PR/CL-040

        - 2021014 & TDPL/PR/CL-003

      - Contact number application engineer:

      - Location: Central Lab, Begumpet Hyderabad.

- Procedure

- Before processing the samples, biosafety cabinet should be cleaned with 70% ethanol.

- Switch ON the UV light for 10- 15min.

- Switch OFF the UV light.

- Switch ON the AIR and Visible light switch situated on the control panel on the right-hand side of the biosafety cabinet.

- Start the spirit lamp and carry out the work.

- After completion of the process switch OFF the Airflow and Spirit lamp.

- Clean any remaining water or waste liquid material spilled on the biosafety platform properly with dry cotton.

- Clean the BSC properly from the ceiling, followed by side glass and then the platform of the chamber with 70% ethanol.

- Spray sterile 70% ethanol and switch OFF the visible light and Switch ON the UV light for 10-15mins.

- Turn OFF the main power.

- Note: Halogenated materials (e.g., bleach) are not recommended for use on stainless steel surfaces of the BSC.

- Maintenance

| Daily Maintenance: | Clean the work area with spirit and cotton. Check if the manometer reading is below 25 mmHg. Check for any residual LPG left in the rubber tube after closing the gas cylinder. Put on the UV light for 30 minutes at the end of the day for decontamination after cleaning. |
| --- | --- |
| Weekly maintenance: | Check the efficacy of UV light by sterility testing of Biosafety Hood. Close the lid and switch on the UV light for 15 minutes. Switch off the UV light. Failure to clean the ultraviolet lamps in the BSC will reduce the lamps’ effectiveness. Periodically clean the ultraviolet lamps in the BSC with a lint-free cloth dampened with alcohol. Air sampling by exposure plates Mark the petri dish plates at the base with the following details with marker: Lot number, Expiry date, Date of exposure. Place the petri dish plate with lid open, exposing the media in the biosafety cabinet and operate the instrument as per SOP given for bio safety cabinet. The plates are kept for 30 mins in the biosafety cabinet and then incubated in an inverted position at 370 C for 24 hours in an incubator. After incubation observe the plates and count the number of colonies in the light source with the help of marker. Acceptance criteria - Less than 20 CFU/Plate In case more than 20 CFU/plate are isolated, then the cleaning, sterilization and sterility check procedure are repeated. Take swabs from the floor, sides, and ceiling inside the working area of hood. Culture them on Blood Agar plate. Incubate at 37oC for 24 hrs. and look for any growth. A negative sterility check of the Biosafety hood indicates the efficacy of UV light |

- Troubleshooting

      - Check the manometer reading.

      - Check the light switch for any loose connection.

      - If all the above are working properly and still the machine is not working, inform the Biomedical Engineering Department to arrange an instrument engineer-visit.

- References

- N/A