SOP – 004

PROCEDURE FOR

OPERATING THE UNI CELL DxI 800

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Test Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	9

9.	Patient Preparation	9

10.	Equipment & Reagents	9

11.	Safety	10

12.	Calibration Procedure	10

13.	Test Procedure	12

14.	Quality Control Procedure	12

15.	Interferences & Cross Reactions	15

16.	Result Calculation	15

17.	Biological Reference Interval	15

18.	Reportable Interval	15

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	16

23.	References	16

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. K. Anusha Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | AMR | Analytical Measurement Range |
| 2 | BC | Biochemistry |
| 3 | °C | Degree Centigrade |
| 4 | CLIA | Chemiluminescence Immunoassay |
| 5 | CV | Coefficient of Variation |
| 6 | CRR | Clinical Reportable Range |
| 7 | EQAS | External Quality Assurance Scheme |
| 8 | HOD | Head of Department |
| 9 | IFCC | International Federation for Clinical Chemistry |
| 10 | IQC | Internal Quality Control |
| 11 | LLOQ | Lower Limit of Quantitation |
| 12 | MSP | Management System Procedure |
| 13 | N/A | Not Applicable/Available |
| 14 | pH | Potential of Hydrogen |
| 15 | QC | Quality Control |
| 16 | QM | Quality Manager |
| 17 | QNS | Quantity Not Sufficient |
| 18 | SD | Standard Deviation |
| 19 | SOP | Standard Operating Procedure |
| 20 | SST | Serum Separator Tube |
| 21 | TRF | Test Requisition Form |
| 22 | TDPL | TRUSTlab Diagnostics Private Limited |
| 23 | ULOQ | Upper Limit of Quantitation |
| 24 | VB12 | Vitamin B12 |

- Objective

- The purpose of this SOP is to provide an overview of the UniCel DxI 800 instrument, its uses, major hardware and software components as well as basic instructions on its operation.

- Scope

- This procedure applies to all staff of the Biochemistry Department tasked with the duty of performing these tests.

- Responsibilities

The Quality Manager is responsible for implementing, monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definition

- Immunoassay: An immunoassay is a biochemical test that measures the presence or concentration of a substance (analyte) in a sample, typically using the binding of antibodies to antigens. It relies on the specificity of the antigen-antibody interaction to detect and quantify the target molecule. Immunoassays are widely used in clinical diagnostics, research, and drug discovery.

- Principle and Method:

The UniCel DxI Access Immunoassay System is an in vitro diagnostic device used for the quantitative, semi-quantitative, or qualitative determination of various analyte concentrations found in human body fluids.

Access Immunoassay System assays are heterogeneous, enzyme-mediated assays. During sample processing, coated paramagnetic particles react with patient sample and assay-specific reagents to form immune complexes. A magnetic field immobilizes the bound immune complexes while the unbound components are washed away. Chemiluminescent substrates react with an enzyme label present in the bound complexes and releases light. A luminometer measures the light emitted, and the system converts this reading, measured in relative light units (RLUs), into a test result.

- Specific Performance Characteristics

The UniCel DxI instrument performance characteristics include instrument operating and temperatures, luminometer performance

Instrument Operating Characteristics: The UniCel DxI system has the following operating characteristics:

      - Warm up

      - Upper cabinet temperature:    The system will not report results when the upper cabinet temperature is outside the acceptable limits of 64°F to 97°F (18°C to 36°C).

      - Processing speed - Maximum throughput rate (UniCel DxI 800 instrument):

          - 400 test results per hour for one-step assays

          - 200 test results per hour for two-step assays

      - LLOQ: 50 pg/ml

      - ULOQ: 1500 pg/ml

Instrument Temperature: The UniCel DxI instrument keeps temperature-sensitive devices and reagents within the following temperature ranges:

      - Sample wheel			4.00°C to 10.00°C

      - Reagent storage			4.00°C to 10.00°C

      - Wash carousel			36.80°C to 37.20°C

      - Substrate (in the incubator)		36.90°C to 37.90°C

Luminometer Performance: The luminometer performs within the following read precision and drift limits:

      - Read precision within a single RV	≤ 0.2% CV

      - Drift					≤ 0.5% in 4 weeks

- Type of Sample

- Serum

- Type of Container & Additives

- N/A.

- Patient Preparation

- N/A.

- Equipment & Reagents

Equipment

- Beckman Coulter DxI 800 CLIA Analyzer

Reagents

- Beckman Reagents

- Safety

Open and close instrument covers to avoid causing personal injury or damage to the instrument.

Never wear exposed metal that can come into contact with electrical components of the instrument.

You will come in contact with potentially infectious materials during these procedures. Handle and dispose of biohazard materials according to proper laboratory procedures. Proper hand, eye, and facial protection is required.

If you have difficulty locating the components listed in any instrument cover procedure, contact your technical support representative.

- Calibration Procedure

- The UniCel DxI system supports three types of assay calibrations:

      - Quantitative

      - Semi-quantitative

      - Qualitative

- The calibrations differ in the way they calculate and report test results.

| Quantitative | In general, calibrator test results provide a multi-point calibration curve. The system uses the calibration curve to convert a measured response in RLUs to an analyte concentration and then expresses the result in numerical units. |
| --- | --- |
| Semi-Quantitative | Calibrator test results provide a multi-point calibration curve. The system uses the calibration curve to convert a measured response in RLUs to an analyte concentration and then expresses the result in numerical units. These assays may report their quantitative result as a qualitative interpretation, such as positive, negative, or equivocal. |
| Qualitative | Calibrator test results provide a cutoff value based on a formula defined in the APF. The system compares a test RLU value to the cutoff value and then classifies the result as reactive or nonreactive for the analyte. In the case of a reactive test result, a confirmatory test may be indicated and automatically requested as a reflex test. The confirmatory assay produces a result of confirmed or not confirmed. |

Calibration Test Request

- Sample Manager:

      - New Request F3  Calibration F2  [Select the Calibrator Lot number]  Ok

- Calibration F1  [Select the Calibrator Lot number]  Enter the Rack ID  Enter

- (Repeat as needed).

      - Place the samples in the racks.

      - Exit the Test Requests screen.

Calibration Failure Codes

| Failure Code | Description |
| --- | --- |
| Bad Fit | The curve does not meet the acceptance criteria defined in the APF. Possible causes are poor precision or a curve that was too flat or too steep. |
| CV Std 0 | The %CV of the S0 (zero) calibrator replicates do not meet the limits defined by the APF |
| Insuff Data | The system did not have enough data to perform the necessary calibration calculations. This failure occurs when two or more replicates are not calculated, due to a short sample or an instrument error. |
| Limits | The RLUs of either a replicate or the mean of the replicates are outside the ranges defined in the APF. |
| Max Iter | The system was unable to create a curve before performing the maximum number of calculation iterations. The system attempts to calculate the curve 100 times before generating this error. |
| No Fit | The system was unable to fit a curve to the data points. Possible causes include math rule violations, such as division by zero or calculating the log of a negative number. |
| Resp Delta | The total response of the curve, which is the difference between the highest RLU reading and the lowest RLU reading, is too small. |
| Temp Out | The upper cabinet temperature changed by more than 6ºC while the system processed the curve replicates. |

- Procedure

The system starts sample processing when it receives a test request either from the laboratory information system (LIS), or from the system console.

Each patient test request consists of:

      - Sample ID

      - Rack number

      - Rack position

      - Sample type

      - Test list

Test requests may contain additional information, including:

      - Patient ID

      - Sample dilution

      - Priority (STAT)

      - Comments

      - Patient demographic information (from LIMS only)

Once a test request is entered, you place the sample in a rack and load it in the sample presentation unit (SPU) onload area. The instrument automatically moves the rack to the sample aliquot station, scans the bar codes on the rack and sample tube, calculates the volume of sample necessary to complete all requested tests and replicates and to meet reserve volume requirements, and aliquots the sample into one or more reaction vessels (RVs) to be stored in the refrigerated sample storage area. The rack is then moved to the SPU offload area, where you can unload and reclaim any unused quantity of sample for further testing.

Sample Processing LIMS Requests:

| Step | Action |
| --- | --- |
| 1 | Create a test request in LIMS for six (6) practice sample(s). Have the LIMS print a barcode for each request. |
| 2 | Apply the barcode to a primary sample tube and add fluid for the practice sample. Note: Pooled serum or control material may be used for practice. |
| 3 | Load the samples into the appropriately designated racks for the sample container(s) available for each practice sample. |
| 4 | Load the racks onto the Sample Processing Unit (SPU). |

      - If the LIMS system prohibits you from following the instructions in the Step/Action table above, use the alternate method below:

| Step | Action |
| --- | --- |
| 1 | Obtain patient samples that are requested in your LIS and require processing. |
| 2 | Load the samples into the appropriately designated racks for the sample container(s) available for each sample. |
| 3 | Load the racks onto the Sample Processing Unit (SPU). |

Instrument Shut Down/Restart:

- PC Admin F7  System Admin F8 Shut Down  Instrument F2  Shut Down Instrument Software [Enter the System Password]  OK F1

- When the software has shut down, a message will display instructing you to turn off the instrument power. Then:

      - Open the lower right door and turn the instrument power switch to the off (O) position.

      - Wait at least 20 seconds, then turn the instrument power switch on.

Maintenance:

    - Utility Routine

      - The Utility Routine is an automated maintenance procedure that runs automatically every 4 hours if the system is not processing samples. Enable the Utility Assay on the instrument.

      - Software navigational path:

- Main Menu-F9 or Menu  Maintenance Review-F5  Enable Utility Routine-F6

    - Daily Maintenance:

      - Back up the system (verify back up successful)

      - Record the test count

      - Shake down the solid waste container

      - Run a clean routine

    - 5,000 &10,000 Test Interval Maintenance:

- The 10,000 test interval maintenance procedures include a step to perform the 5,000 test interval maintenance. Perform daily maintenance procedures after you complete 10,000 and 5,000 test interval procedures.

      - Perform 10,000 and 5,000 test interval maintenance. Use the references    in the Help to ensure all procedures are completed. For efficiency, perform the procedures in the order listed below:

- Reference: from the Maintenance Review screen Help

- Maintenance Overview

- Test Interval Maintenance: 10,000 Tests

- Use the green link for each procedure

        - Install clean aspirate probes

        - Replace the duck bill valve

        - Remove and inspect dispense probes

        - Inspect:

          - Pump connections at the wash buffer reservoir

          - Upper aspirate peristaltic pump connections

          - Reagent pipettors and clean the wash towers

          - For fallen RVs

          - Liquid waste drawer

          - Wash buffer supply drawer

      - Close the covers and initialize the system.

      - While the system is initializing:

        - Clean the air filter

        - Clean the aspirate probes

      - Daily Maintenance

        - Verify system performance by processing QC.

        - Locate the Maintenance Log in the Instructions for Use Manual. Ask your instructor for a practice log and initial the maintenance performed.

- Quality Control Procedure

- Use the Help System to identify detailed instructions for quality control. Use the following software path to access the detailed instructions. Follow all cautions and warning when performing quality control.

      - Reference: from the Quality Control screen Help

- Setting Up Quality Controls

- Adding or Editing a Quality Control

      - Reference: from the Sample Manager screen Help

- Patient Test Requests

- Running Quality Control Tests

- Entering QC Test Requests

      - Reference: from the Quality Control screen Help

- Quality Control Overview

- QC Status Indicators and Flags

      - Follow the steps in the Step/Action table below

| Step | Action |
| --- | --- |
| 1 | Create a QC test request for the Practice QC1 and Practice QC 2 samples, requesting tests for which mean(s) and SD(s) were defined in Step 1 of the Practice Exercise. |
| 2 | Pour the appropriate QC material into a 2 ml cup programmed for Practice QC1. |
| 3 | Pour water into a 2 ml cup programmed for Practice QC2. |
| 4 | Complete the steps listed in the flowchart in the “Quality Control” tab of the workbook to process the QC that was just requested. |

- Interferences & Cross Reactions

- N/A.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A

- Potential Source of Variability

- N/A

- References

- System Operation Manual