SOP – 008

PROCEDURE FOR

OPERATING BIO-RAD D10

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Procedural Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	8

13.	Procedure	8

14.	Quality Control Procedure	14

15.	Interferences & Cross Reactions	15

16.	Result Calculation	15

17.	Biological Reference Interval	15

18.	Reportable Interval	15

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | HOD | Head of Department |
| 5 | HBV | Hepatitis B Virus |
| 6 | HIV | Human Immunodeficiency Syndrome |
| 7 | HM | Haematology |
| 8 | IQC | Internal Quality Control |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | NS | Normal Saline |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | Rpm | Revolutions per minute |
| 16 | SOP | Standard Operating Procedure |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to describe the procedure for operating the Bio-Rad D10 HPLC instrument, in a safe and proper manner. This SOP also provides information about the procedural principles and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definitions

- HPLC: High-performance liquid chromatography, formerly referred to as high-pressure liquid chromatography, is a technique in analytical chemistry used to separate, identify, and quantify specific components in mixtures.

- Procedural Principle

        - .

- Specific Performance Characteristics

- N/A

- Type of Sample

- Blood collected following universal procedure for venipuncture.

- Type of Container & Additives

- N/A.

- Patient Preparation

- N/A.

- Equipment & Reagents

Equipment: BIO-RAD D10

- Safety

N/A.

- Calibration Procedure

- N/A.

- Procedure

Maintenance

    - DAILY MAINTENANCE

      - Pre-Run Checklist

        - Check Method (Lot Info)

        - Check Buffers and Wash Solution Levels

        - Reset Number of Injections if New Bottle(s) Installed

        - Confirm correct line goes to the appropriate bottles

        - Check Cartridge Injection Count (Lot Info)

        - Reset Number of Injections if New Cartridge Installed

        - Check External Waste Tank

        - Check Pump Pressure (Maintain/Start Pump)

        - Flowrate should be same as during analysis (1.5ml/min)

        - Check for Leaks During Pressure Check

        - Check Paper Supply

      - Post-Run Checklist

        - Remove Samples

        - Dispose of Primer, Calibrator and Controls in Biohazard Container

        - Wipe Spills

    - MONTHLY MAINTENANCE

      - Checklist

        - Clean Exterior and Interior Surfaces - Wipe Down with Deionized Water

        - Clean and Decontaminate the Sample Path

        - Clean the Dilution Chamber

        - Clean and Inspect the Sample Racks - Use Water

        - Clean the Internal Waste Bottle.

    - Troubleshooting

      - Pressure Problems - Check for:

        - Low Pressure

        - Leaks in the Fluid System

        - Dummy Cartridge left in Cartridge Holder

        - Air Bubbles in the Pumps

        - High Pressure

        - Overtightened Fittings

        - Crimped Tubing

        - Clogs in the Fluid System - Cartridge most common

        - May also be caused by blocked Injector

      - Chromatography (High area counts) – Check for:

        - High Patient Hct

        - Inadequate Wash of Sample Probe

        - Inadequate Drain of Dilution Chamber

        - Syringe Pump not Functioning Properly

        - Unmixed Pre-diluted Sample

      - Chromatography (High area counts) – Check for:

        - Short Sample - Check Sample Volume

        - Clots in Sample

        - Low Patient Hct

        - Air in Sample Lines or Syringe

        - Leak at Dilution Chamber.

      - Retention Time – Check for:

        - Temperature Incorrect (Not Adjustable)

        - Buffer Contamination or Evaporation

        - Incompatible Buffer Lots

        - Wrong Cartridge Lot for Buffer Lot

        - Defective Cartridge

        - Slow Leak in Flow Path

        - Air Bubbles

        - Faulty Column Heater

        - Pump Issue - Broken Piston, Bad Seal

        - Faulty Proportioning Valve

        - Degasser removing too much dissolved gas- split peak (saturated O2 is stripped from the Hgb molecule)

        - Wrong Diskette

- Quality Control Procedure

| Quality Control Management | 3 controls levels (low, normal, high) Target values download (USB) Levey-Jennings Graphs Radar Graphs XB on 3 or 9 parameters, mean of 20 runs Compatible with QCP |
| --- | --- |

- Interferences & Cross Reactions

- N/A.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

- D-10 Hemoglobin Testing System Operation Manual