SOP – 009

PROCEDURE FOR

OPERATING HORIBA YUMIZEN H500

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definition	7

5.	Procedural Principle	7

6.	Specific Performance Characteristics	7

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	8

13.	Procedure	8

14.	Quality Control Procedure	14

15.	Interferences & Cross Reactions	15

16.	Result Calculation	15

17.	Biological Reference Interval	15

18.	Reportable Interval	15

19.	Procedure when Results are not in Range	15

20.	Critical Alert Values	15

21.	Laboratory Clinical Interpretation	15

22.	Potential Source of Variability	15

23.	References	15

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Abdul Aleem Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | EDTA | Ethylenediamine Tetra-acetic Acid |
| 3 | EQAS | External Quality Assurance Scheme |
| 4 | HOD | Head of Department |
| 5 | HBV | Hepatitis B Virus |
| 6 | HIV | Human Immunodeficiency Syndrome |
| 7 | HM | Haematology |
| 8 | IQC | Internal Quality Control |
| 9 | MSP | Management System Procedure |
| 10 | N/A | Not Applicable/Available |
| 11 | NS | Normal Saline |
| 12 | QC | Quality Control |
| 13 | QM | Quality Manager |
| 14 | QNS | Quantity Not Sufficient |
| 15 | Rpm | Revolutions per minute |
| 16 | SOP | Standard Operating Procedure |
| 17 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to describe the procedure for operating the Automated Haematology Analyzer – Horiba Yumizen H500, in a safe and proper manner. This SOP also provides information about the procedural principles and other details.

- Scope

- This procedure applies to all staff of the Haematology Department tasked with the duty of performing these tests.

- Responsibilities

It is the responsibility of the technicians of the Haematology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff involved in this process (including contract personnel) are responsible for complying with this procedure.

- Definitions

- Impedance: This is the core method for counting cell populations. When a blood sample is passed through an aperture, the cells disrupt the electrical current, and the change in resistance is measured. This allows for counting and classifying cells based on their size and volume.

- Procedural Principle

The HORIBA Yumizen H500 hematology analyzer primarily uses the impedance principle for counting and enumerating red blood cells, white blood cells, and platelets. It also utilizes photometric methods, specifically absorbance, to measure the concentration of substances in the blood sample. Additionally, the H500 can perform a 6-part differential count of white blood cells using its iDoubleDiff technology.

- Specific Performance Characteristics

- N/A

- Type of Sample

- Blood collected following universal procedure for venipuncture.

- Type of Container & Additives

- N/A.

- Patient Preparation

- N/A.

- Equipment & Reagents

Equipment: Horiba Yumizen H500

- Safety

N/A.

- Calibration Procedure

- N/A.

- Procedure

Start of Day

      - To Check the Waste Container Level:

        - Check the level of waste in the container.

        - If it needs to be emptied, refer to:

- Maintenance and Troubleshooting > Replacement Procedures > To Replace Reagents >To Replace the Waste Container

To Perform a Manual Startup

      - Select Startup

      - Wait until the cycle is over.

      - A startup cycle takes approximately one minute.

      - Blank cycles (cycles without any blood specimen) are performed during the startup cycle. The startup passes if the background counts are within acceptable limits as given in the following table:

| - Parameter | - Background Count Limits |
| --- | --- |
| - WBC | - ≤ 0.3 x 109/l |
| - RBC | - ≤ 0.03 x 1012/l |
| - HGB | - ≤ 3 g/l |
| - PLT | - ≤ 5 x 109/l |

To Check Control Results

- Access: Main Screen > Quality Assurance > Quality control

      - Select a Control lot.

      - Check that the results are within control target values range.

      - If results are out of range, perform a concentrated cleaning and rerun the Control blood.

      - NOTE: If results are still out of range, please check the reagents and the stability of the Control blood. If these are okay, then contact the local HORIBA representative.

Running Blood Samples

      - To Create an Order:

        - Press Stat mode in the function toolbar.

        - Enter the sample ID if necessary.

        - Since the instrument is interfaced with TRUSTlab LIMS, press Query to retrieve the analysis data for the sample.

        - In the event that the instrument is unable to connect to LIMS:

          - Enter the physician name, the department and/or comments about the sample, if necessary.

          - If necessary, enter the patient ID or press Search PID to select it.

      - To Run a Blood Sample

- Make sure that the order for the blood sample has been created and validated or that it is received from LIMS.

        - Gently and thoroughly mix the sample.

        - Open the tube and place it below the sampling needle. Lift it so that the needle can sample its content.

        - Press the Sampling bar.

- Remove the tube and put the cap back on once the needle has moved up.

      - Out of Linearity Range Alarm:

        - The “Out of linearity range” alarm is generated when the results are out of the minimum and maximum values for linearity or visibility ranges defined for the instrument.

        - The sample may be diluted with ABX Diluent and rerun.

| Parameter | Triggered if… | Consequence |
| --- | --- | --- |
| HGB | Result out of linearity range | The HGB, MCH and MCHC results are flagged with an "*" |
|  | Result out of visibility range | The HGB, MCH and MCHC results are rejected |
| WBC | Result out of linearity range | WBC, LYM#, MON#, EOS#, BAS#, NEU# and LIC# results are flagged with an "*" |
|  | Result out of visibility range | WBC, LYM#, MON#, EOS#, BAS#, NEU# and LIC# results are rejected |
| RBC | Result out of linearity range | RBC and HCT results are flagged with an "*" |
|  | Result out of visibility range | RBC and HCT results are rejected. MCV, MCH, MCHC, RDW-CV and RDW-SD results are flagged with an "*" |
| PLT | Result out of linearity range | RBC, PLT, PCT, PDW, MPV, P-LCC and WBC results are flagged with an "*" |
|  | Result out of visibility range | and WBC results are rejected. HCT, MCV, MCH, MCHC, RDW-CV and RDW-SD results are flagged with an "*" |

To Check the Waste Container Level

      - Check the level of waste in the container.

      - If it needs to be emptied, refer to:

- Maintenance and Troubleshooting > Replacement Procedures > Replacing Reagents > To Replace the Waste Container

      - When disposing of waste, protective clothing must be worn (lab coat, gloves, eye protection, etc.).

      - Follow your local and/or national guidelines for biohazard waste disposal.

      - At the beginning of each day, before startup, check if the waste container needs to be emptied.

      - During instrument operation, do not remove the reagent tubing and the liquid waste tubing under any circumstance.

      - Waste must be handled according to your local and/or national regulations.

Starting the Instrument

- To Switch the Instrument On

      - Before switching the instrument on, you need to:

        - Check the operational conditions described in the

          - Introduction > Operational Conditions chapter.

        - Check all instrument connections. To learn more about connections, refer to the

          - Introduction > Labels and Connections chapter.

        - Check if the waste container needs to be emptied. Follow the instructions in the

          - Specifications > Reagent Specifications > Waste Handling Precautions chapter

      - Switch the instrument on.

      - Wait during initialization. A startup cycle begins, if scheduled by default.

      - Log in to the application. Refer to the

          - Workflow > Start of Day > Starting the Instrument > To Log in to the Application chapter

To Log in to the Application

- NOTE: If an error message is displayed during initialization or if the application does not start properly, please contact the local HORIBA representative.

      - Select your username.

      - Enter your password.

      - Press Validate in the contextual toolbar.

      - Select Reset sample ID auto-numbering in the Start of Day window, if necessary. Note: This step is only necessary at the beginning of the day, and if the option is selected.

      - Select Erase work list in the Start of Day window, if necessary. This step is only necessary at the beginning of the day and if the option is selected.

      - Select Export patient results before data deletion if necessary. This option is available if 10000 results are detected in the system. You can select the XML or PDF format to export your data.

          - Memory	10,000 results + Graphics

      - Press Validate

- If a reagent has expired, the software informs you when you log in. If the shutdown cycle has not been performed at the end of the previous day, the system forces you to perform a shutdown cycle before the startup. The shutdown cycle is efficient and valid only if the cleaner remains at least 10 minutes in the chambers after the cycle. This cleans the hydraulic circuit. You must not perform any action during these 10 minutes and risk the repeat of the shutdown

To Perform a Manual Startup

      - Press Startup

- Wait until the cycle is over. A startup cycle takes approximately one minute. Blank cycles (cycles without any blood specimen) are performed during the startup cycle. The startup passes if the background counts are within acceptable limits.

| Parameter Count limits | Background |
| --- | --- |
| WBC | ≤ 0.3 x 109 /L |
| RBC | ≤ 0.03 x |
| HGB | ≤ 3 g/L |
| PLT | ≤ 5 x 109 /L |

      - You can consult the startup results in the Blank cycle logs area.

To Run a QC in Stat Mode

      - Prepare your Control blood according to the specific instructions detailed in the control blood package insert.

      - Press Stat mode in the function toolbar.

      - Enter the sample ID of the Control blood sample.

      - Press Validate in the contextual toolbar.

      - Gently and thoroughly mix the sample.

      - Put the tube in the tube holder and close the door.

      - When placing your tube in the tube holder, make sure that you are using the appropriate position and that the cap can be pierced. If not, remove the cap before closing the door.

      - Press Validate to start sampling.

      - When the tube holder moves back to the front of the analyzer, remove the tube and put the cap back on if needed.

      - Risk of erroneous results if the specimen is not continuously mixed between each analysis. Keep on mixing the specimen between each analysis.

To Check Control Results

      - Select a Control lot.

      - Make sure that the results are within control target values range.

      - If results are out of range, perform a root cause analysis and corrective action.

      - If results are still out of range, please check your reagents and your Control blood stability and then contact your local HORIBA representative.

To Manage the Reagents

      - The system can manage HORIBA Medical reagents automatically (levels and expiration date). It informs the user about the reagents status at the end of the instrument start, or displays an alarm message in the Reagents screen if a reagent runs low or has expired. However, it is recommended to check the reagent levels and expiration date before starting the system. To do so, you need to:

        - Check the level of the reagent bottles from the software.

        - Visually check the lot number and expiration date on the reagent bottles.

        - If a reagent bottle has to be changed, refer to the

          - Maintenance and Troubleshooting > Replacement Procedures > Replacing Reagents chapter

Other Important Features

| Data management | Linux™ Transmit patient & QC to LIS Uni-directional or bi-directional connection Communication Protocol: ASTM |
| --- | --- |
| Logs | Reagents, calibration, Quality Control, blank cycle, users, settings, service, communication |
| Parameters (27) | WBC, NEU# & NEU%, LYM# & LYM%, MON# & MON%, EOS# & EOS%, BAS# & BAS%, LIC# & LIC%*, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD*, PLT, MPV, PCT*, PDW*, P-LCC*, P-LCR* |
| Sample-Volume | CBC mode: 20 µL DIFF mode: 20 µL |
| Stability | 48 hours post-draw stability for CBC parameters, 24 hours post-draw stability for DIFF parameters |
| Operating sample mode | Open or Closed Tube |

- Quality Control Procedure

| Quality Control Management | 3 controls levels (low, normal, high) Target values download (USB) Levey-Jennings Graphs Radar Graphs XB on 3 or 9 parameters, mean of 20 runs Compatible with QCP |
| --- | --- |

- Interferences & Cross Reactions

- N/A.

- Result Calculation

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Critical Alert Values

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

- Horiba Operator Manual