SOP – 013

PROCEDURE FOR

COOLING CENTRIFUGE

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Principle	7

5.	Parts of the Instrument	7

6.	Required Equipment’s & Reagents	8

7.	Environment & Safety Control	8

8.	Calibration Procedure	9

9.	Equipment Record	9

10.	Quality Control Procedure	10

11.	Procedure	10

12.	Maintenance	10

13.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | DW | Distilled water |
| 3 | HOD | Head of Department |
| 4 | MSP | Management System Procedure |
| 5 | N/A | Not Applicable/Available |
| 6 | PPE | Personal Protection Equipment |
| 7 | QC | Quality Control |
| 8 | QM | Quality Manager |
| 9 | rpm | Revolutions per minute |
| 10 | SOP | Standard Operating Procedure |
| 11 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this SOP is to provide guidelines for operating and maintaining the equipment in the Molecular Biology Department, to ensure safe and proper usage.

- Scope

- This procedure applies to all staff of the Molecular Biology Department.

- Responsibilities

It is the responsibility of the technicians of the Molecular Biology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff that may be involved in this process (including contract personnel) will be responsible for complying with this procedure.

- Principle

The Centrifuge 5424 R is used for the separation of aqueous solutions and suspensions of different densities in approved sample tubes.

- Parts of the Instrument

      - Centrifuge lid

      - Monitoring glass: Visual control for rotor stop or speed control option using stroboscope

      - Display

      - Control panel: Keys and dials (dependent on the device version) for operating the centrifuge.

      - Emergency release

      - Nameplate

      - Mains/power switch: Switch for switching the centrifuge on and off.

      - Mains/power connection: Connection socket for the mains cable supplied.

      - Fuse holder

      - Interface for software updates: Only for authorized service personnel

      - Condensation water tray

- Required Equipment’s & Reagents

      - 70% ethanol

      - Tissue

      - PPE

      - Disposable gloves

      - Eye Protection

      - Clean water & detergent

- Environment & Safety Control

| Warning for intended use | Electric shock due to damage to the device or main cord Lethal voltage inside the device Danger due to incorrect voltage supply Risk of injury when opening or closing the centrifuge lid Risk to injury from rotating rotor Damage to device due to spilled liquids Danger from open tube lids Damage to health due to limited aerosol tightness of incorrect use. |
| --- | --- |

| Depiction | Meaning | Location |
| --- | --- | --- |
|  | Observe the safety instruction in the operating manual | Right side of the device |
|  | Observe operating manual | Right side of the device |
|  | Warning: possible hand injury | Upper side of the device |
|  | Always tight the rotor with enclosed rotor key | Upper side of the device under the centrifuge lid |
|  | Warning of biological risk when handling infection liquids or pathogenic germs | Aerosol- tight fixed angle rotor: rotor lid |

- Calibration Procedure

Frequency: Once a year.

A centrifuge calibration consists of three parameters, speed, time and temperature.

      - Time:

          - Set the time foe 5min using set parameters& record result, repeat the test for ten times.

      - Calibration of rpm (speed)

          - Fixed rotor

          - Set the rpm of rotor to 13000 & record the result by tachometer.

          - Set rpm of rotor to 6000 & record the result.

          - Swing out rotor.

          - Set the rpm of rotor to 4000 & record the result.

          - Set rpm of the rotor to 2000 & record the result.

      - Calibration of temperature:

          - Fill the two-sample tube with DW & place it in the sample holder. Set the parameter like temp 5 C at 13000rpm & record the result by using digital thermometer.

          - Fill the two-sample tube with DW & place it in the sample holder. Set the parameter like temp 10 C at 13000rpm & record the result by using digital thermometer.

          - Fill the two-sample tube with DW & place it in the sample holder. Set the parameter like temp 20 C at 13000rpm & record the result by using digital thermometer.

- Equipment Record

      - Equipment Name: Cooling centrifuge

      - Manufacturer Name: Eppendorf Centrifuge 5424R

      - Serial No & equipment ID: 5404JL943173 & TDPL/PR/CL-030

      - Contact number application engineer:

      - Location: Central Lab, Begumpet Hyderabad.

- Quality Control Procedure

- N/A

- Procedure

- Switch the centrifuge on at the mains/power switch.

- The parameter settings of the last run are displayed.

- The lid opens.

- Place the sample or reagent in the rotor by proper balancing.

- Pushdownthecentrifugeliduntilthelidlatchengagesandthelidisautomaticallyclosed.

- The centrifuge will close automatically.

- Note: The Open key lights up blue.

- As per the protocol set the rpm, time & temperature of the centrifuge.

- Maintenance

| Cleaning and disinfecting the device: | Open the lid. Switch the device off at the mains/power switch. Disconnect the mains/power plug from the voltage supply. Loosen the rotor nut by turning the rotor key counterclockwise. Remove the rotor. Clean and disinfect all accessible surfaces on the device including the mains/power cord using a damp cloth and recommended cleaning agents. Thoroughly clean the rubber seals of the rotor chamber with water. Rub the dry rubber seals with glycerin or talcum powder to prevent them from becoming brittle. Other components of the device, such as the lid latch, motor shaft and rotor cone, must not be lubricated. Clean the motor shaft with a soft, dry, lint-free cloth. Do not grease the motor shaft. Check the motor shaft for damage. Check the device for corrosion and damage. Leave the centrifuge lid open when the device is not used. Only connect the device to the power supply if it is fully dry inside and out. |
| --- | --- |

- References

- N/A