SOP – 015

PROCEDURE FOR

LAMINAR AIR FLOW

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Principle	7

5.	Required Equipment’s & Reagents	8

6.	Environment & Safety Control	8

7.	Calibration	9

8.	Equipment Record	9

9.	Quality Control Procedure	9

10.	Procedure	10

11.	Maintenance	10

12.	References	11

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | HEPA | High Efficiency Particulate Air |
| 3 | HOD | Head of Department |
| 4 | LAF | Laminar Air Flow |
| 5 | MSP | Management System Procedure |
| 6 | N/A | Not Applicable/Available |
| 7 | PPE | Personal Protection Equipment |
| 8 | QC | Quality Control |
| 9 | QM | Quality Manager |
| 10 | SOP | Standard Operating Procedure |
| 11 | TDPL | TRUSTlab Diagnostics Private Limited |
| 12 | ULPA | Ultra-Low Particulate Air |
| 13 | UV | Ultraviolet |

- Objective

- The purpose of this SOP is to provide guidelines for operating and maintaining the Laminar Air Flow in the Molecular Biology Department, to ensure safe and proper usage.

- Scope

- This procedure applies to all staff of the Molecular Biology Department.

- Responsibilities

It is the responsibility of the technicians of the Molecular Biology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff that may be involved in this process (including contract personnel) will be responsible for complying with this procedure.

- Principle

A laminar air flow workstation is a closed cabinet fitted with HEPA filtered air flow system with unidirectional constant flow of air with almost no or minimal turbulence. The purpose of using such workstations in laboratory is to create particle and bacteria free working environment to carry out specialized work. As these unit’s discharge air towards user, they provide no personal protection but product protection from room contaminants.

Types:

VERTICAL LAMINAR AIR FLOW: It is the most used type of laminar air flow. In the vertical flow cabinet, room air is entered in working area through HEPA filter fitted on top of the cabinet. Thus, air flows downward (vertically) towards work surface and leaves the cabinet sweeping out particles and bacteria.

HORIZONTAL LAMINAR AIR FLOW: In the horizontal air flow cabinet, room air is entered in working area through HEPA filter fitted on the back side of the cabinet. Filtered air flows horizontally towards the user creating

- Required Equipment’s & Reagents

      - 70% ethanol

      - Tissue

      - PPE

      - Disposable gloves

      - Eye Protection

      - Clean water & detergent

      - Sodium Hypochlorite

- Environment & Safety Control

      - NEVER OPERATE UV LAMP WITHOUT EYE AND SKIN PROTECTION

      - Any internal adjustment, modification or maintenance to this equipment must be undertaken by qualified service personnel.

      - The use of any hazardous material in the cabinet must be monitored by an industrial hygienist, safety officer or some other suitably qualified individual.

      - Explosive or inflammable substances should never be used in the cabinet unless a qualified safety professional has evaluated the risk involved.

      - decontamination measures are to be applied for the workstation and accessories,

      - measures are to be taken in the case of an accident.

      - Repairs to the device must be carried out only by trained and authorized expert personnel.

      - Keep these operating instructions close to the unit so that safety instructions and important information are always accessible.

| Safety alert symbol indicates a potentially hazardous situation which, if not avoided, could result in death of serious injury |  |
| --- | --- |
| Safety alert symbol indicates a potentially hazardous situation which, if not avoided, may result in minor or moderate injury. |  |
| CAUTION used without the safety alert symbol indicates a potentially hazardous situation which, if not avoided, may result in property damage. | CAUTION: |
| Used for important information | NOTE: |
| Biohazard |  |
| Chemical Hazard |  |

- Calibration

      - Frequency: Once a year.

      - The complete range of testing services to ensure that laminar flow equipment provides the necessary clean air for your work processes.

        - Airflow Velocity and Uniformity Test

        - HEPA/ULPA Filter Installation Leak Test

        - Particle Counting Test

        - Noise Level Test

        - Light Intensity Test

      - In addition, the following secondary maintenance Services required:

        - Supply and Installation of Pre-filters and HEPA Filters

        - Supply and Installation of Blower

        - Repair of Stainless Steel - Electrical/Mechanical Safety Checks

- Equipment Record

      - Equipment Name: Laminar Air flow

      - Manufacturer Name: ICLEAN (Integrated Cleanroom Technologies Pvt. Ltd.)

      - Serial No & equipment ID:

        - 2021025 & TDPL/PR/CL-041

        - 2021026 & TDPL/PR/CL-042

        - N/A & TDPL/PR/CL-033

      - Contact number application engineer:

      - Location: Central Lab, Begumpet Hyderabad

- Quality Control Procedure

- N/A

- Procedure

| Preparation 250 mL of 70% Ethanol: | Add 175 mL of Absolute Ethanol (99.9%) in 75mL Distilled Water |
| --- | --- |
| Preparation of 1000 mL of 1% Sodium Hypochlorite: | Add 200 mL of 5% NaOCl to 800 mL of Distilled water. |

- Before processing the samples, LAF (vertical) should be cleaned by 70% ethanol.

- Switch ON the UV light for 10- 15min.

- Switch OFF the UV light.

- Switch ON the AIR and Visible light switch situated on the control panel on the right-hand side of the LAF.

- Start the spirit lamp and carry out the work.

- After completion of the process switch OFF the Airflow and Spirit lamp.

- Clean any remaining water or waste liquid material spilled on the laminar airflow platform properly with dry cotton.

- Clean the LAF properly from the ceiling, followed by side glass and then the platform of the chamber with 70% ethanol.

- Spray sterile 70% ethanol and switch OFF the visible light and Switch ON the UV light for 10-15mins.

- Maintenance

- Follow the suggested maintenance schedule in order to maintain your cabinet at its optimum performance.

| Daily: | Thoroughly surface-decontaminate the work surface, inner back walls, and interior surface of the window using 70% ethanol or other disinfectant depending on the user's requirement. Note: Do not use any disinfectant containing chlorine-based substance as this will cause corrosion to steel and stainless steel resulting in irreparable damage to the cabinet structure. Do surface decontamination as well to the UV lights. When cleaning the interior, the operator should not insert any part of his/her body into the work zone (except hands). |
| --- | --- |
| Monthly | Using a damp cloth, clean the exterior surfaces of the cabinet, particularly the front and top of the cabinet, to remove any accumulated dust. Check all the service fixtures (when present) for proper operation. |
| Annually | Have the cabinet recertified by a qualified certification technician. Replace the UV lamp. |

- References

- N/A.