SOP – 017

PROCEDURE FOR

OPERATING THE AUTOCLAVE

- TABLE OF CONTENTS

TABLE OF CONTENTS	2

APPROVAL AUTHORITY	3

RECORD OF AMENDMENTS	4

DISTRIBUTION LIST	5

LIST OF ABBREVIATIONS	6

1.	Objective	7

2.	Scope	7

3.	Responsibilities	7

4.	Definitions	7

5.	Procedural Principle	7

6.	Specific Performance Characteristics	8

7.	Type of Sample	8

8.	Type of Container & Additives	8

9.	Patient Preparation	8

10.	Equipment & Reagents	8

11.	Safety	8

12.	Calibration Procedure	9

13.	Procedure of Instrument Operation	9

14.	Quality Control Procedure	11

15.	Interferences & Cross Reactions	11

16.	Result Calculation	11

17.	Biological Reference Interval	11

18.	Reportable Interval	11

19.	Procedure when Results are not in Range	11

20.	Critical Alert Values	11

21.	Laboratory Clinical Interpretation	11

22.	Potential Source of Variability	11

23.	References	11

Flowchart: Procedure for Handling Biosafety Cabinet	12

- APPROVAL AUTHORITY

| Prepared by: | Head of Department | Dr. Sadiya Shahnaz Name | Date: 14/11/2024 |
| --- | --- | --- | --- |
|  |  | Signature |  |
| Reviewed & Approved by: | Laboratory Head | Dr. Abdul Aleem Name | Date: 14/11/2024 |
|  |  | Signature |  |
| Issued by: | BME Manager | Pallavi Elisetty Name | Date: 14/11/2024 |
|  |  | Signature |  |

- RECORD OF AMENDMENTS

| S. No. | Ver. No. | Ver. Date | Rev. No. | Rev. Date | Description of Changes |
| --- | --- | --- | --- | --- | --- |

| Document Copy No. | Name / Designation of the holder of the document copy | Type of Copy | Version No. | Revision No. |
| --- | --- | --- | --- | --- |
| 01/10 | Managing Director | Controlled Copy | 01 | 00 |
| 02/10 |  |  |  |  |
| 03/10 |  |  |  |  |
| 04/10 |  |  |  |  |
| 05/10 |  |  |  |  |
| 06/10 |  |  |  |  |
| 07/10 |  |  |  |  |
| 08/10 |  |  |  |  |
| 09/10 |  |  |  |  |
| 10/10 |  |  |  |  |

- DISTRIBUTION LIST

- LIST OF ABBREVIATIONS

| S. No. | Abbreviations | Expansion |
| --- | --- | --- |
| 1 | °C | Degree Centigrade |
| 2 | HOD | Head of Department |
| 3 | MSP | Management System Procedure |
| 4 | N/A | Not Applicable/Available |
| 5 | QC | Quality Control |
| 6 | QM | Quality Manager |
| 7 | SOP | Standard Operating Procedure |
| 8 | TDPL | TRUSTlab Diagnostics Private Limited |

- Objective

- The purpose of this standard operating procedure (SOP) is to ensure that all potentially infectious waste materials, and waste that must be made biologically inactive before disposal, are adequately sterilized when subjected to autoclaving. This SOP will outline procedures for waste collection, treatment by autoclaving, validating autoclave performance, and a protocol for managing an autoclave that is not operating correctly (if a positive biological indicator is generated during testing of performance).

- Scope

- This procedure applies to all staff of the Microbiology Department. This SOP applies to all autoclaves that are used to decontaminate infectious waste materials at TRUSTlab.

- Responsibilities

It is the responsibility of the technicians of the Microbiology Department to comply with this SOP.

HOD shall ensure implementation.

The Quality Manager is responsible for monitoring, reviewing and ensuring compliance with this procedure.

All other TRUSTlab staff that may be involved in this process (including contract personnel) will be responsible for complying with this procedure.

- Definitions

      - Potentially infectious waste, which requires autoclaving, includes the following:

      - Biological waste: includes blood and blood products, excretions, exudates, secretions, and other body fluids that cannot be directly discarded into the municipal sewer system, and waste materials saturated with blood or body fluids.

      - Culture and stocks: Include etiologic agents of disease and associated biologicals, including specimen cultures and dishes and devices used to transfer, inoculate and mix cultures, wastes from production of biologicals, and serums and discarded live and attenuated vaccines. Also includes gloves and other disposable personal protective equipment used as barriers when handling biological wastes.

- Procedural Principle

Moist heat sterilization using autoclave is commonly used for the sterilization of biohazardous trash, heat and moisture resistant materials such as aqueous preparation (culture media).  This method is also used for the sterilization of surgical dressings and medical devices.

The most common type of steam sterilizer in the microbiology laboratory is the gravity displacement type. Another type of autoclave is Vacuum/Gravity Assisted.

Moist heat destroys microorganisms by the irreversible denaturation of enzymes and structural proteins. The temperature at which denaturation occurs varies inversely with the amount of water present. Sterilization in saturated steam thus requires precise control of time, temperature, and pressure.

Pressure serves as a means to obtain the high temperatures necessary to quickly kill microorganisms. Specific temperatures must be obtained to ensure microbicidal activity. Minimum sterilization time should be measured from the moment when all the materials to be sterilized have reached the required temperature throughout.

The usual procedure is to heat at 1.1 kg/cm2 [15 pounds/square inch (lb/in2)] steam pressure, which yields a temperature of 121°C. At 121°C, the time of autoclaving to achieve sterilization is generally considered to be 15-60 minutes, depending on the volume of the load.  To ensure that sterilization is successful one should ascertain that:

      - The air in the chamber is completely evacuated so that the chamber fills with steam.

      - Articles are placed in a distributed manner in the autoclave so that steam can easily penetrate them.

- Specific Performance Characteristics

- N/A

- Type of Sample

- N/A.

- Type of Container & Additives

- N/A.

- Patient Preparation

- N/A.

- Equipment & Reagents

Equipment:

- Autoclave

Reagents & Consumables

      - Signaloc label

      - Distilled water

      - Cotton

      - Tissue

      - Autoclavable bags

- Safety

Autoclave (vertical) should be installed vertically on a level floor, in a well-ventilated area, away from the main working area.

Spills must be reported and cleaned up immediately.

Never use excessive force when sealing or opening the chamber door.

The earth wire (base copper) is provided outside the boiler cover. The connection to the extension-wire should be soldered. The earthing connection should be made of a proper earth electrode/point.

Pressure switches once adjusted for required pressure should not be tampered or meddled with, unnecessarily.

If the pressure inside the chamber does not reach 15 lb/in2 the sterilization is incomplete and can be dangerous to use. Hence, it is essential that pressure reaches the optimum point. Otherwise, re-sterilization is recommended.

Do not permit crowding of packs inside the sterilizing chamber into tight masses, as this may lead to inadequate exposure of all surfaces to the sterilizing steam.

Neutralize waste containing bleach with equal amounts of 1% sodium thiosulfate in water prior to autoclaving.

The autoclave door should not be opened until the chamber temperature falls below 100°C

If steam shoots out from any part of the chamber door (or any other part of the machine other than the drainpipes), push the OFF button immediately. If the steam was leaking from the edge of the door, wait for the chamber to cool below 100°C then you may try to reseal the door. If the problem persists, leave the machine off and report the problem.

- Calibration Procedure

For Time:

      - Set the timer on the autoclave for a full cycle. Hold the stopwatch and be prepared to click it.

      - Turn on the autoclave and the stopwatch simultaneously.

      - Repeat the process if the time and the stopwatch do not match.

For Temperature:

      - Place the maximum registering thermometer near the built-in temperature sensor.

      - Load and operate the autoclave as usual.

      - Record the maximum autoclave temperature according to the built-in temperature display.

      - Let the autoclave cool.

      - Remove the maximum registering thermometer from the autoclave and record the temperature reading in the logbook.

      - Record the discrepancy if the maximum temperature of the built-in display is different from the maximum temperature on the registering thermometer.

      - Post a correction factor on the autoclave control panel if the temperatures were different.

For Pressure:

      - Calibrate the timer and temperature gauges before these steps.

      - Load and operate the autoclave as usual. When the temperature display shows 121 degrees Celsius, the pressure gauge should show 15 pounds per square inch (15 psi).

      - Record the discrepancy, if any.

      - Mark the correct point for 15 psi on the cover of the gauge and post a correction factor on the autoclave if this is not true.

- Procedure of Instrument Operation

Start Up Procedure

      - Pack all the materials that must be autoclaved.

      - Stick the autoclave label (Signaloc label) to any one of the containers to be autoclaved.

      - Check the water level in the autoclave. If necessary, add or remove water. It should be filled up to 1/4th of the level.

      - See the heating element is submerged in water.

      - Load the autoclave with all the necessary materials to be autoclaved.

      - Close the lid and screw tightly.

      - Check if the steam valve is open.

      - Switch on the autoclave and record the time. Wait until the steam is released.

      - Close the steam valve and record the time.

      - Wait until the pressure reaches 15 psi and record the time.

      - For sterilization – leave the autoclave on for 15 minutes at 15 psi pressure.

      - For decontamination - leave the autoclave on for 60 minutes at 15 psi pressure.

Shutdown Procedure

      - Open the steam valve.

      - Switch off the machine and record the time.

      - Wait until the pressure reaches ‘0’ psi and open the autoclave.

      - Check for the autoclave label color. (This should turn from blue to dark grey/black upon effective autoclaving.)

      - Remove all the autoclaved materials.

      - Decant the water and rinse. Close the lid.

      - Use fresh water every day.

      - Wash the autoclave with soap and water once a week.

Maintenance

    - Weekly Maintenance:

      - Clean with soap powder and wash with tap water.

      - After rinsing with water, refill with fresh water for reuse.

Troubleshooting

      - Check if the steam valve is working properly.

      - Check if the pressure gauge shows the prescribed pressure of 15 psi in working condition.

      - Check for leakage of steam anywhere in the instrument.

      - Check the gasket.

      - If all the above are working properly and still the machine is not working, inform the Biomedical Engineering Department to arrange for a service engineer’s visit.

- Quality Control Procedure

- N/A.

- Interferences & Cross Reactions

- N/A.

- Biological Reference Interval

- N/A.

- Reportable Interval

- N/A.

- Procedure when Results are not in Range

- N/A.

- Laboratory Clinical Interpretation

- N/A.

- Potential Source of Variability

- N/A.

- References

- N/A

Before processing the samples, the biosafety cabinet should be cleaned by 70% ethanol.

Switch ON the UV light for 30 min.

Switch OFF the UV light. Wait for 5 minutes

Switch ON the blower & Visible light switch situated on the control panel on the right-hand side of the biosafety cabinet.

Start the spirit lamp and carry out the work

After completion of the process switch OFF the blower and Spirit lamp

Clean any remaining water or waste liquid material spilled on the biosafety platform properly with dry cotton.

Clean the BSC properly from the ceiling, followed by side glass and then the platform of the chamber with 70% ethanol

Spray sterile 70% ethanol and switch OFF the visible light and Switch ON the UV light for 10-15mins

Turn OFF the UV light and the main power.

Flowchart: Procedure for Handling Biosafety Cabinet